## Slide 1
Scheduled Calibration of
IMS Infrasound Stations
Yacine Sid Ahmed, Benoit Doury,
Alfred Kramer, Ichrak Ketata
CTBTO Preparatory Commission, Vienna, Austria
Presentation Date: 20 June 2023, 12:00
PTS-680

## Slide 2
Quality Assurance Processes1
PTS-680
1 Quality Assurance Infrastructure for IMS Infrasound Measurements (ECS/WGB-52/PTS/8)

## Slide 3
Operational Manual Requirements1
PTS-680
1 [CTBT/WGB/TL-11,17/17/Rev.7]
“Stations are provided with the means to carry out a calibration whose result can be compared against a reference that was established at the time the station was certified or when a revalidation was completed at the station”
The calibration performed at the time of the station certification or revalidation is referred to as “initial calibration”. “The result of the initial calibration forms the baseline for certification and future calibrations”
The calibration performed on a regular basis to verify that the system response remains “within tolerances” of the baseline is referred to as “on-site calibration”.
“When the results of the calibration are not within tolerances, the station operator informs the Technical Secretariat and initiates the required maintenance.”
On-site calibrations are performed at least on a yearly basis
Both initial and on-site calibrations are “full frequency response calibrations”
System baseline
sensitivity
+ response
5% threshold
1 year

## Slide 4
Principle: side-by-side comparison
PTS-680
1 In situ calibration of atmospheric-infrasound sensors including the effects of wind-noise-reduction pipe systems
2 Quality Assurance Infrastructure for IMS Infrasound Measurements (ECS/WGB-52/PTS/8)
A passive calibration method described by Gabrielson1 was chosen2 for on-site calibration
Station equipped with calibration capabilities:
  reference system: sensor + pipe + inlet port
  second data stream forwarded t

## Slide 5
Principle: processing in Vienna
PTS-680
Scheduled Calibration process supported by software (CalxPy) for IS stations equipped with calibration capabilities (results computation, review, results archiving)
Calibration results computation is centralized in Vienna
Calibration is continuous
[see ePoster on CalxPy]

## Slide 6
Scheduled Calibration Process
PTS-680
[ECS-WGB-59-PTS-1]
Responsibility shared between the Station Operator and the PTS.
Processes in 4 phases
PLANNING
IMPLEMENTATION
REVIEW
PUBLICATION
Verification that calibration is possible (via CalxPy, IRS)
Automatic results computation (via CalxPy)
Results reviewed at the PTS (via CalxPy)
Troubleshooting/maintenance initiated if applicable (via IRS)
Validated calibration results prepared for being shared with Authorized Users (CalxPy)
Update of computed sensitivities in the digitizers (followed via IRS, data frames)

## Slide 7
Scheduled Deployment of calibration capabilities - Status
PTS-680
Deployment of calibration capabilities is ongoing at infrasound stations of the IMS network

## Slide 8
Results: overview
PTS-680
32% of the IS network is equipped with calibration capabilities (17 stations; 128 channels)
25% of the IS network have calibration results reported to Member States (13 stations; 99 channels)
[ECS-WGB-59-PTS-1]

## Slide 9
Examples of review: overview of all sites
PTS-680
[see ePoster on CalxPy]
Amplitude
Phase
All sites
IMS passband

## Slide 10
Examples of review: detailed site review
PTS-680
[see ePoster on CalxPy]
1 - Sensitivity tracker @4s
2 - Averaged sensitivity
3 - Averaged response
4 - # coherent windows
5 - Time selection
6 - IMS2.0 result
1
2
3
4
5
6

## Slide 11
Findings
PTS-680
1 Gabrielson T., Quality Effects of Spacing between a Reference Sensor and an IMS Infrasound Element
on the Accuracy of Frequency-Response Estimation
Unexpected sensitivity and unexpected responses are investigated
  Identification of IDC misconfiguration
  Identification of instrument change without configuration update
  Identification of sensors sensitivity shift at specific pressure levels [ITW 2019]
Identify equipment integration issues at stations (optimal installation of reference systems 1)
  change inlet port location at sites
  use of 2 reference inlet ports (case of no primary manifold)

## Slide 12
Principle: processing in Vienna
PTS-680
2 issues identified and fixed
  inverted connection between operational & reference channels at the digitizer level following maintenance activity
  reference sensor DQ issues
1 issue identified (flooding): site relocation to be planned
[see ePoster on CalxPy]

## Slide 13
Principle: processing in Vienna
PTS-680
2 issues identified and fixed
  inverted connection between operational & reference channels at the digitizer level following maintenance activity
  reference sensor DQ issues
1 issue identified (flooding): site relocation to be planned
[see ePoster on CalxPy]

## Slide 14
Principle: processing in Vienna
PTS-680
2 issues identified and fixed
  inverted connection between operational & reference channels at the digitizer level following maintenance activity
  reference sensor DQ issues
1 issue identified (flooding): site relocation to be planned
[see ePoster on CalxPy]

## Slide 15
Principle: Going further with Data Quality Control
PTS-680
Sensitivity trackers for selectable time periods and frequencies
Sensitivity as a function of the absolute pressure
Sensitivity as a function of the temperature
[see ePoster on CalxPy]
Responses over selectable time periods
(predefined, configurable)
Refined responses using specific algorithms
(e.g. identify low spectral levels)

## Slide 16
Benefits
PTS-680
Significant progress in a relatively short time period
Scheduled Calibration Process is operational at the PTS
Success:
  compliance with requirements on calibration
  maintenance actions triggered
  validation that most channels have a stable response which reassures current PTS installation practices

## Slide 17
Perspective
PTS-680
Roll-out calibration capabilities in the IS network
Refine/Develop additional software features (e.g. advanced data QC)
Passive and continuous calibration gives access to a goldmine of data (~2TB)
  How can we get even more out of this dataset?
    Improve station design to further improve stability of systems sensitivity
    (e.g., vault design, sensor models for specific environments, testing of sensors in labs)
     Any suggestion? …
CalxPy will be made available shortly in NIAB to authorized users
For further information, see ePosters on the software tool [P3.1-578]
and the calibration process of IMS stations [P4.1-714].

## Slide 18
Perspective
PTS-680
Thank you for your attention