## Slide 1: Determination of the frequency response of seismic & infrasonic IMS station sensors using an on-site calibration approach
BGR- B4.3 – Federal Seismological Survey, CTBT (German NDC)
Michaela Schwardt1, Christoph Pilger1, Thomas Bruns2, Leonard Klaus2, Samuel Kristoffersen3, Franck Larsonnier3, Lars Ceranna1, Patrick Hupe1, Peter Gaebler1
1Federal Institute für Geosciences and Natural Resources (BGR)
2Physikalisch-Technische Bundesanstalt (PTB)
3 Commissariat à l’énergie atomique et aux énergies alternatives (CEA)
O3.1-698

## Slide 2: Need for on-site calibration procedures
Missing standards & calibration procedures
In the low-frequency range down to 0.01 Hz reliable calibration procedures, which include traceability to SI, are currently missing  rely on the manufacturer's specifications
Joint Research Project 19ENV03 Infra-AUV*
Development of primary & secondary calibration methods
Establish procedures, which allow permanent on-site calibration without any interruptions of the recordings
Consideration of traceability & measurement uncertainties
* https://www.ptb.de/empir2020/infra-auv/home/
O3.1-698

## Slide 3: Need for on-site calibration procedures
Installation of a laboratory calibrated instruments as reference sensor
X
multi-component calibration system of PTB
Laboratory
Seismometer in a vault
Field Sensors
?
Interruption of measurements
Violation of technical requirements
If the field sensor will not come to the laboratory,
the laboratory calibration will go to the field sensor. (freely adapted from Francis Bacon, “Of Boldness”, 1625)
O3.1-698

## Slide 4: Development of on-site calibration methods
Existing calibration methods are “relative”
They rely on sensor comparison using a common excitation signal & calculate a gain ratio between the sensors to characterize "errors" in the transfer function (e.g., Pavlis & Vernon, 1994; Sleeman et al., 2006)
Application of a modified approach of Gabrielson (Gabrielson, 2011; Charbit et al., 2015; Green et al., 2021)
  Determination of the gain ratio between a co-located reference sensor (REF) and station sensor under test (SUT) and inclusion of various similarity measures (e.g., coherence, cross-correlation)
SUT & REF measure the same coherent signal
Known response of REF
(Amplitude & Phase)
Ensures traceability to SI
O3.1-698

## Slide 5: Algorithm behind the on-site calibration approach
*1 Ensure stationarity of time series (Charbit et al., 2015)
*2 Cross-Array coherency (Green et al., 2021)
Advantages?
The filtering & the choice of the window length allows a better determination of the response function within the individual frequency ranges
Longer data segments with larger windows enable analysis of the low frequency range
*Similarity measures:
Magnitude-squared cohereny (MSC)
Cross-correlation
Cross-Array coherency*2
Segment i
Division into segments
9 Windows, 50% overlap
Filter*1 within Pass-bands
Calculations of similarity measures & cross-/auto-spectral densities
Raw data signal
Similarity measure* ≥ Threshold?
Determination of gain ratio between REF & SUT
Calculation of frequency response function (amplitude & phase) of SUT using traceable response of REF
O3.1-698

## Slide 6
Laboratory calibration of 2 vertical station seismometers (GS13) and a 3-component seismometer (STS2.5) by PTB
Full frequency response is known & traceable to SI; uncertainties are given
Installation of the calibrated seismometers at the station in August 2022
260 days of continuous data have been recorded
Removal of STS2.5 in May 2023 for a laboratory re-calibration
STS2.5
CMG-3T
GS13
STS2.5
CMG-3T & GS13
ca. 2m
Fieldtest for the on-site calibration of seismometers
O3.1-698

## Slide 7
CMG-3T (SUT) & STS2.5 (REF)
Calculation of complex gain ratio on a daily basis
mean over 260 days
Similarity measures taken into account: MSC ≥ 0.98 XC ≥ 0.8
Calc. Amplitude Response
mean gain ratio
0.05 - 8 Hz: low standard deviations
> 8 Hz: high standard deviations, deviation from nominal value (> ±5% of nom. Value)
< 0.05 Hz: high standard deviations
Co-located 3-component seismometer STS2.5 vs. CMG-3T (vertical component)
O3.1-698

## Slide 8
Below 0.05 Hz and above 5 Hz the transfer functions for the horizontal components could only be determined with high standard deviation in the laborator by now! The reference is back in the laboratory for re-calibration!
Co-located 3-component seismometer STS2.5 vs. CMG-3T – Amplitude & Phase
O3.1-698

## Slide 9
Cross-check
Both sensors are laboratory calibrated ↔ verification of the method
! Caution different input gain (1x vs. 8x) & pre-amplifier (40V/V) ↔ effects must be taken into account
Co-located broadband & short-period seismometer GS13 vs. STS2.5 (Z-component)
GS13 (SUT) & STS2.5 (REF)
Similar behavior at high frequencies (>5 Hz) to the 3-component seismometers
O3.1-698

## Slide 10
In a subsequent step, we aim at the calibration of all sensors of a station using a single temporary and stationary reference sensor.
Comparison between 2 laboratory calibrated GS13 seismometer
Distance between sensors: ≈ 1425m
Station-wide seismometer calibration?
O3.1-698
SUT: C2
REF: C7
SUT: C7
REF: C2
< 1 Hz  > 1 Hz  still work to do!

## Slide 11: Summary
Added value
Cross-check shows that the on-site calibration method provides values comparable to those obtained in the laboratory
the results are closer to the true values than those of the electrical calibration (former calibration method)
Open Points
Uncertainty propagation
Cause of deviation for frequencies > 5 Hz
On-Site calibration with traceable calibrated reference allows for a traceable calibration of station sensors without interrupting the measurements!
O3.1-698

## Slide 12: THANK YOU!
Acknowledgements
This project 19ENV03 Infra-AUV has received funding from the EMPIR Programme co-financed by the Participating States and from the European Union’s Horizon 2020 research and innovation programme. We would like to thank everyone at PTB involved in the laboratory calibration of the sensors, as well as everyone from BGR responsible for the installation of the sensors at the stations
References
Charbit, M., Doury, B., & Marty, J. (2015). Evaluation of infrasound in-situ calibration method on a 3-month measurement campaign. In Infrasound Technology Workshop 2015 (ITW2015).
Gabrielson, T. B. (2011). In situ calibration of atmospheric-infrasound sensors including the effects of wind-noise-reduction pipe systems. The Journal of the Acoustical Society of America, 130(3), 1154-1163.
Green, D. N., Nippress, A., Bowers, D., & Selby, N. D. (2021). Identifying suitable time periods for infrasound measurement system response estimation using across-array coherence. Geophysical Journal International, 226(2), 1159-1173.
Schwardt, M., Pilger, C., Gaebler, P., Hupe, P., & Ceranna, L. (2022). Natural and Anthropogenic Sources of Seismic, Hydroacoustic, and Infrasonic Waves: Waveforms and Spectral Characteristics (and Their Applicability for Sensor Calibration). Surveys in Geophysics, 43(5), 1265-1361.
O3.1-698
* https://www.ptb.de/empir2020/infra-auv/home/

## Slide 13
O3.1-698

## Slide 14: Calibration of an Infrasound station including the wind-noise reduction system
REF: MB2005 (single inlet) SUT: MB3a (WNRS, 96 inlets)
Calculation of gain ratio for each day (>80 days), averaging over all days
Similarity measures: Coherency (MSC) >0.98 Cross-Array Coherency >0.6 (Green et al., 2021)
O3.1-698

## Slide 15
In a subsequent step, we aim at the calibration of all sensors of a station using a single temporary and stationary reference sensor.
Comparison between 2 laboratory calibrated GS13 seismometer
Distance between sensors: ≈ 1425m
Station-wide seismometer calibration?
O3.1-698
SUT: C2
REF: C7
SUT: C7
REF: C2
< 1 Hz  > 1 Hz  still work to do!
After averaging: