## Slide 1
Automatic Analysis of Two-Dimension Coincidence Spectra
Hugues Paradis, Charles-Philippe Mano, Philippe Gross, Anthony Der Mesrobian Kabakian
CEA/DAM
CEA/DAM is developing new b/g measurement systems ; in order to improve the analysis methods at low count rates, a new multiplicative spectral unmixing algorithm for full-spectrum analysis was tested on a particular setup for different types of mixtures.
P3.6-886
Spectral unmixing is the process of decomposing a mixed spectrum into several contributing species based on their respective spectral signatures, to which abundances are attributed. Using an iterative process, the abundance of each radionuclide is calculated.
The spectral unmixing algorithm finds reconstructed activities of complex mixtures with good accuracy. Notably, the uncertainties are smaller than for the classical method. Characteristic limits (detection threshold and limit) are improved in many cases.
The first results with this new algorithm show that this method is more accurate with the use of all the information in the spectrum. Characteristic limits are improved in many cases. The next step is the test the algorithm with measured data from stations of the IMS.
Please do not use this space, a QR code will be automatically overlayed

## Slide 2
Introduction : 2D b/g spectrum
P3.6-886
Please do not use this space, a QR code will be automatically overlayed
The Comprehensive nuclear-Test Ban Treaty (CTBT) is an international treaty open to signature in 1996 to prevent any nuclear test. In this context, the CTBT organization deployed the International Monitoring System (IMS) to measure seismic, hydro-acoustic, infrasound and radionuclide signals to detect potential nuclear events. Out of the 336 stations of the IMS, around 40 are dedicated to the measurement of radioxenon (131mXe, 133mXe, 133Xe and 135Xe) in the air.
A detector under developement in the laboratory is composed of a well NaI(Tl) detector for X and g measurement and a silicon detector (PIPSBox) for b measurement. Both these measurements are done in coincidence, and data are analyzed as 2D spectra.
Example of a low β-γ resolution spectrum
ROI 135Xe
ROI 133Xe
ROI 133mXe
Energy NaI(Tl) (keV)
Energy Si (keV)
ROI 131mXe
PMT
NaI(Tl)
PIPSBox
Measurement system

## Slide 3
Hypotheses and objectives
P3.6-886
Please do not use this space, a QR code will be automatically overlayed
Hypothesis :
The global spectrum is the sum of individual spectra (well known)
Each radionuclide has its own spectrum (efficiency is the global counting of the spectrum)
Use of the individual abundance for the quantification
Objectives :
Use of all the information (uncertainties and characteritics limits improved)
Automatic analysis (without expert)
abundance
Energy (channel)
Example for a 1D spectrum

## Slide 4
Method : spectral unmixing
P3.6-886
Please do not use this space, a QR code will be automatically overlayed
214Bi
214Pb
133Xe
135Xe
133mXe
Probability to have the canal i implemented if there is a decay of the RN
Probability a posteriori to have a decay of RN if a canal i is incremented
Probability a priori to have a decay of the RN
Advantages :
Maximized counting efficiency
Validated method on 1D spectra (H Paradis et al., Appl. Radiat. Isot. 158 (2020) 109068 ; R André et al., Metrologia 58 (2021) 015011)
All the information of the spectrum is used
Limitations :
High computational resources needed for SD/LD (parallelization on a GPU computing server)
131mXe
Spectra for each radionuclide of interest individually

## Slide 5
Results : accuracy and characteristic limits
P3.6-886
Please do not use this space, a QR code will be automatically overlayed
The algorithm is compared with the classical ROI method. Trueness, accuracy and characteristic limits are compared.
Spectral unmixing reconstructs activities fairly accurately, with uncertainties smaller than with the ROI method.
The DT and DL are lower for the spectral unmixing algorithm, except for 135Xe.
Radionuclide | Expected counting | A simulated (mBq) | AROI method (mBq) | Aspectral unmixing (mBq)
214Bi | 238 | 40 | 40.7 6.2 | 38.9 4.9
214Pb | 316 | 40 | 41.2 5.7 | 39.3 3.4
135Xe | 1666 | 171 | 170.5 3.8 | 169.9 3.4
133Xe | 51 | 20 | 17.8 13.5 | 21.0 6.2
133mXe | 79 | 28 | 28.5 3.4 | 28.0 2.5
131mXe | 12 | 6 | 5.8 3.4 | 5.6 1.8
Radionuclide | Expected counting | A simulated (mBq) | AROI method (mBq) | Aspectral unmixing (mBq)
214Bi | 238 | 40 |  | 
214Pb | 316 | 40 |  | 
135Xe | 1666 | 171 |  | 
133Xe | 51 | 20 |  | 
133mXe | 79 | 28 |  | 
131mXe | 12 | 6 |  | 
 | Detection threshold (DT) |  |  |  | Detection limits (DL) |  |  | 
Radionuclide | DTROI (count) | DTunmix (count) | DTROI (mBq) | DTunmix (mBq) | DLROI (count) | DLunmix (count) | DLROI (mBq) | DLunmix (mBq)
135Xe | 8.7 | 12.8 | 1.2 | 1.3 | 20.1 | 31.3 | 2.7 | 3.22
133Xe | 13.2 | 3.4 | 10.4 | 1.3 | 29.2 | 10.7 | 23.0 | 4.12
133mXe | 3.3 | 1.0 | 1.6 | 0.3 | 9.3 | 4.0 | 4.6 | 1.43
131mXe | 3.3 | 1.2 | 1.6 | 0.6 | 9.3 | 4.0 | 4.6 | 1.85
Example of an analysis of a mix with 6 radionuclides (and background). Activities and uncertainties obtained with 1000 different spectra analysed
Detection thresholds and detection limits with 300 hits of background in the spectrum (1000 different spectra analysed)

## Slide 6
Results : characteristics limits
P3.6-886
Please do not use this space, a QR code will be automatically overlayed
The spectral unmixing algorithm gives non Gaussian distributions for the DT and DL :
    Depending of the radionuclide
    Calculated by Monte-Carlo simulations
It is necessary to determine them before the start of any measurement

## Slide 7
Results : characteristics limits
P3.6-886
Please do not use this space, a QR code will be automatically overlayed
In many cases, a small amount of Rn can be present in the measurement cell
→ Daughters (214Pb and 214Bi) are detected in b/g measurement
→ signal overlay
→ Increase of DT and DL
The DT is lower for the spectral unmixing algorithm in all cases
The DL is lower for the spectral unmixing algorithm in all cases, except for 135Xe
Evolution of DT with the activity of 222Rn
Evolution of DL with the activity of 222Rn

## Slide 8
Conclusion
P3.6-886
Please do not use this space, a QR code will be automatically overlayed
Development of a spectral unmixing algorithm based on the global analysis of the spectrum
Activities reconstructed with a good accuracy
Uncertainties lower than with the classical method (ROI)
Detection threshold and detection limits globally better than with ROI method on background spectra
 Improvement for 135Xe :
    Reduction of energy cut in b channel (100 keV → 30 keV)
    Optimization on spectral signatures (cut of the high energies to limit mix with 214Bi)
 Analysis of spectra from the IMS
What’s next?