## Slide 1
Comparison of 127Xe Measurements
Troy Robinson,1 Matthew Goodwin,2 Matthew Watrous1
1Idaho National Laboratory (USA), 2Atomic Weapons Establishment (UK)
127Xe can interfere with Nuclear Explosion Monitoring causing false positive hits.
Training detection systems to recognize and quantify 127Xe may reduce this problem.
Px.x-xxx
127Xe was produced by neutron irradiation and quantified using HPGe and beta-gamma detectors.
Attenuation correction was developed for the HPGe analysis.
Results between INL (HPGe) and AWE (β-γ) show good agreement
INL
0.998 ± 0.040 (1σ)
AWE
0.967 ± 0.099 (1σ)
Methods of quantification of 127Xe were demonstrated through a laboratory intercomparison exercise.
Recommendations for implementation of 127Xe quantification in the IMS are suggested
Please do not use this space, a QR code will be automatically overlayed

## Slide 2
Introduction (1 of 2)
127Xe interference in Nuclear Explosion Monitoring
T3.2-779
Please do not use this space, a QR code will be automatically overlayed
Non-fission product xenon radionuclides were recently observed during air-sampling equipment testing in Tennessee, USA1
  Radionuclides identified were: 122Xe, 125Xe, and 127Xe
Presence of these non-fission product xenon radionuclides resulted in a significant number of false positive detects of traditionally monitored fission product xenon radionuclides (131mXe, 133Xe, 133mXe, 135Xe)
  127Xe contributes coincidence counts in the ROIs for fission product radionuclides2
Prudence dictates that samplers monitoring the traditional four radionuclides would benefit from training to identify the presence of other xenon radionuclides such as 127Xe to avoid reporting false positive detections
Fig 1. 127Xe counts overlap the 7-ROIs used for detection of fission product xenon radionuclides. Figure is from Klingberg et al.2

## Slide 3: Introduction (2 of 2) Challenges to 127Xe Detection
127Xe (along with 125Xe and 129mXe) have never been detected in 20 years of IMS monitoring3
  Detecting these radionuclides is unlikely to occur without more sensitive systems3
    Xenon International and SAUNA-III systems offer improved sensitivity that may lead to more detections
Training systems to detect 127Xe is challenging due to the following:
  Calibration sources3
  Memory effect for non-coated beta cells
  Increased uncertainty in all radionuclides under analysis because of the inclusion of a larger number of interference terms3
T3.2-779

## Slide 4
Objectives (1 of 1)
Please do not use this space, a QR code will be automatically overlayed
Produce 127Xe by neutron irradiation
Develop/employ purification of 127Xe from 125I by-product
Perform 127Xe detector calibration
Develop self-attenuation correction for 127Xe quantification
Develop container that enables sampling gas both by syringe or by expansion
Quantitatively dilute and share sample with a partner laboratory for an intercomparison test
T3.2-779

## Slide 5
Methods/Data (1 of 4)
127Xe Production
Please do not use this space, a QR code will be automatically overlayed
T3.2-779
Isotopically-enriched 126Xe obtained from ISOFLEX USA (San Francisco, USA)
Two quartz ampoule were sealed with 0.5 cm3 of 126Xe
Ampoules were analyzed by XRF to verify integrity of ampoules
Ampoules were irradiated by Washington State University (USA) 10 MW Triga reactor
Isotope | 124Xe | 126Xe | 128Xe | Other Xe
Content (%) | 0.024 | 99.94 | 0.014 | 0.003
Isotopic Enrichment of 126Xe gas (from ISOFLEX Certificate)

## Slide 6: Methods/Data (2 of 4) 127Xe Purification
T3.2-779
Irradiation of 126Xe results in production of two isotopes:
  127Xe (T1/2 = 36.346 days)
  125Xe (T1/2 = 16.9 hr)
Targets were aged for 10 days to allow decay of 125Xe  γ + 125I
Radioiodine is separated from radioxenon using a dry-ice/ethanol slurry cooled trap
  Iodine is retained and xenon flows through the trap
Gamma Spectra of received 127Xe. 125Xe is present after 24 hours of decay, but absent following 240 hours of decay.

## Slide 7: Methods/Data (3 of 4) 127Xe Quantification
T3.2-779
High Purity Germanium (HPGe) well detectors calibrated using a simulated gas multi-energy gamma standard (Eckert and Ziegler, USA) in a borosilicate flame sealed ampoule geometry at 24cm
127Xe isolated in a borosilicate flame sealed ampoule was quantified at 24cm
The quantified 127Xe was then used as a calibration source at geometries closer than 24cm (12cm, 4cm, 0cm and in the well) to obtain detector efficiencies
  Some of these closer geometries are expected to experience cascade and random summing
  These summing losses are not corrected for, but are wrapped up in the observed detector efficiency
24cm
12cm
4cm
0cm
Well
HPGe

## Slide 8: Methods/Data (4 of 4) Well Detector Summing in Gamma Spectroscopy
T3.2-779
High Purity Germanium (HPGe) well detectors experience summing due to:
  High detector geometric efficiency
  127Xe coincident gamma and x-ray emissions
Peak Type | Energy (keV) | Intensity (%)
X ray (X1) | 28.3, 28.6 | 25.2, 46.5
X ray (X2) | 32.2, 32.3, 33.0 | 4.33, 8.35, 2.47
Gamma (γ1) | 145.25 | 4.31
Gamma (γ2) | 172.13 | 25.7
Gamma (γ3) | 202.86 | 68.7
Gamma (γ4) | 374.99 | 17.3
Decay Emission Data for 127Xe4

## Slide 9
Results (1 of 4)
Surprising Attenuation Correction Findings
Please do not use this space, a QR code will be automatically overlayed
T3.2-779
An experiment to determine the effect of added xenon gas on self-attenuation of the detected counts was performed
A ~0.002 cm3 sample of 127Xe was assayed
The sample was then mixed with ~2.0 cm3 of natural xenon and re-assayed to determine the attenuation effect
The addition of 2.0 cm3 and assay sample was repeated until the total sample volume was equal to just over 8.0 cm3 total volume
An increase in count rate was observed as more xenon was added to the sample
  Added xenon was attenuating x-rays that caused a reduction in summing losses in the gamma ray energy ROIs

## Slide 10: Results (2 of 4) New Shipping Container
T3.2-779
A new shipping container was employed with hybrid sampling capability:
  Conventional manual valve closure to tank
  Expansion Tee with septum for gas syringe sampling

## Slide 11: Results (3 of 4) Beta-Gamma Quantification Development
A seven ROI method was used for quantifying 127Xe in the beta—gamma system
A variety of gas compositions (xenon gas volume ranged from 1-7 cm3; Volume quantified by GC) with a total gas pressure of just below 1 bar were used for assays
A decrease in efficiency was observed as total xenon volume increased
  ROI 3 and 4 dropped by approximately 65% over the range
Because of the overlap of ROI 1 with the ROIs used in quantification of 133mXe and 131mXe, it may be possible to calculate the efficiency of 127Xe from the 133mXe and 131mXe efficiencies (without calibrating the system).
  Perhaps this lends itself to a bootstrap calibration method
  This could also be used as a QA check gas for 133mXe and 131mXe calibration that has a shelf-life
T3.2-779

## Slide 12: Results (4 of 4) Comparison Results
T3.2-779
A sample of 127Xe gas was prepared at Idaho National Laboratory (USA) and shipped to Atomic Weapons Establishment (UK) using the new hybrid sampling shipping container
Target concentration for this sample was 1.0 Bq/cm3
INL performed the analysis using HPGe detectors in the “bean” borosilicate counting geometry
AWE performed the analysis using beta-gamma coincidence counting
Laboratory | Reported Concentration (Bq/cm3)
INL | 0.998 ± 0.040 (1σ)
AWE | 0.967 ± 0.099 (1σ)

## Slide 13
Conclusion (1 of 2)
Summary
Please do not use this space, a QR code will be automatically overlayed
T3.2-779
A sample of 127Xe gas was prepared at Idaho National Laboratory (USA) and shipped to Atomic Weapons Establishment (UK)
Overlap between 127Xe beta-gamma spectrum with 133mXe and 131mXe could enable:
  a use of 127Xe as a QA check material that has a shelf-life for 133mXe and 131mXe
  Known 133mXe and 131mXe could be used in a boot-strap calibration for 127Xe quantification without ever calibrating the system.
Results show good agreement and indicate that 127Xe is a radionuclide that can be used for intercomparison exercises
Radioxenon monitoring systems would benefit from being able to identify 127Xe when present
  Although not a fission product, 127Xe can be present in the environment as an activation product
  127Xe if present creates false positive hits for all 4 xenon radionuclides of interest (131mXe, 133Xe, 133mXe, 135Xe)

## Slide 14: Conclusion (2 of 2) Recommendations for 127Xe Analysis in IMS
In agreement with Ely et al.3 it is recommended that higher sensitivity radioxenon monitoring systems be trained to recognize 127Xe presence in a sample (Xenon International, SAUNA III or similar capable systems). Legacy systems likely don’t have the sensitivity to measure atmospheric levels of 127Xe.
Due to the potential introduction of excessive uncertainty in the analysis of the four xenon radionuclides of interest3 (131mXe, 133Xe, 133mXe, 135Xe), 127Xe should not be incorporated in the analysis unless it is detected. Positive hits for the 4 xenon radionuclides of interest should be screened for the presence of 127Xe.
Only systems that have coated cells should undergo 127Xe calibration and training to avoid extended long-lived 127Xe residue that would interfere with normal monitoring operations.
T3.2-779

## Slide 15
References (1 of 1)
Please do not use this space, a QR code will be automatically overlayed
T3.2-779
Eslinger et al. “Determining the source of unusual xenon isotopes in samples” J Environ Radioact (2022) 247:106853.
Klingberg et al. “127Xe coincidence decay analysis in support of CTBT verification” J Radioanal Nucl Chem (2015) 305:225-232.
Ely et al. “First observations of environmental 125Xe, 127Xe, and 129mXe” CTBTO SnT 2021 O2.4-138 (Oral presentation).
NuDat 2.8 obtained through https://www.nndc.bnl.gov/nudat3/chartNuc.jsp (accessed on 4/28/2023). [Nuclear Data Sheets 112, 1647 (2011).