## Slide 1
P3.1-860
Temperature Dependence of Frequency Response
of Short Period Seismometers
N. Ackerley and Z. Gias
Canadian Hazards Information Service, Natural Resources Canada
Broadband calibration and pole-zero fitting can be routine.
Single-frequency calibrations at YKA have long been affected by temperature. It is impossible to meet CTBTO ±5% specification unless this is accounted for.
Short-period sensor damping and sensitivity are temperature-sensitive (but not resonant frequency).
Published response of Yellowknife short-period array sensors needs correction.
There is an unexplained pole in the YKA response near 65 Hz.
Chart: Temperature [°C]
| -36 | -35.69 | -34.44 | -34.119999999999997 | -33.81 | -33.5 | -33.19 | -32.25 | -31.62 | -31 | -28.81 | -25.69 | -21.31 | -20.059999999999999 | -15.69 | 5 | 10.56 | 17.440000000000001 | 26 | 28.62 | 30.19
freq01_hz | -1.9803 | -1.9636 | -2.077 | -2.115 | -2.0887 | -2.0906 | -2.1335 | -2.2293 | -2.1943 | -2.2012 | -2.5164 | -2.0439 | -1.0582 | -1.0244 | -1.0999 | -0.64599 | -0.69749 | -1.0825 | -0.96999 | -0.89129 | -0.91419
damping01 | 19.2821 | 19.3757 | 19.5353 | 19.3803 | 19.3402 | 19.5006 | 19.4969 | 19.5519 | 19.3878 | 19.2217 | 18.6232 | 17.6172 | 15.7524 | 15.7498 | 14.7984 | 10.3545 | 9.8818 | 8.955 | 7.5254 | 6.2348 | 6.5952
freq2_hz | 0.32881 | 0.31463 | -0.23343 | 0.042985 | 0.098956 | -0.21493 | -0.071493 | -0.20523 | -0.082538 | -0.24717 | -0.33299 | -0.46179 | -0.87493 | -1.1094 | -1.1706 | -2.162 | -2.2446 | -2.4611 | -3.1332 | -2.7991 | -2.9242
gain | 11.107 | 11.162 | 10.8514 | 10.7823 | 10.8948 | 10.8424 | 10.7657 | 10.6379 | 10.551 | 10.3088 | 9.1782 | 8.6739 | 7.8207 | 7.6418 | 6.5253 | 1.707 | 1.2367 | -0.14635 | -2.5539 | -3.485 | -3.1367
%/°C | model | obs. | R^2 | comment
gain | -0.223 | -0.221 | 0.998 | excellent agreement
damping01 | -0.143 | -0.206 | 0.990 | same sign & order of magnitude
freq01_hz | -0.017 | 0.022 | 0.684 | poor fit, poor agreement
freq2_hz |  | -0.048 | 0.973 | minimally temperature-sensitive
Original fit:
Additionally fit
pole near 65 Hz:
?