## Slide 1
A. Couture1, A. Carman1, H. Miley2, H. Garland1, L. Lidey1, M. Larche1, M. Moore1, N. Elmore1, R. Guerrero1, S. Cheng1, and T. Butcher1
1Pacific Northwest National Laboratory
2Desert Research Institute
Presentation Date: 9 September 2025
A Gamma-Gamma Coincidence Detection System for Radionuclide Aerosol Monitoring
O3.2-251

## Slide 2
Motivation
Current IMS radionuclide aerosol monitoring stations collect, decay, and count – 24 hrs each (3 days total)
Filter collection & HPGe measurement – either manual or automated
Recent atmospheric transport studies indicate improved detection and location performance of the IMS network with shorter sampling intervals
Objective: Maintain IMS requirements & shorten collection/counting intervals
Proposed Solution: Gamma-gamma coincidence detection & ESP collection (RASA 2.0)
A Gamma-Gamma Coincidence Detection System for Radionuclide Aerosol Monitoring
A. Couture, A. Carmen, H. Miley, H. Garland, L. Lidey, M. Larche, M. Moore, N. Elmore, R. Guerrero, S. Cheng, and T. Butcher
O3.2-251
PW Eslinger, et al. (2023) Journal of Env. Rad., Vol 257, 107088, https://doi.org/10.1016/j.jenvrad.2022.10708
IMS Detection Improvement 24-hr to 12-hr Sampling

## Slide 3
Timing, Flow, and Detector Efficiency
Shorter collection times
  → Better ATM transport
  → Increases MDC
Rules for Timing
  Count = Collect (required)
  Count + Decay + Collect =72 hrs (required?)
Extended decay times lower 140Ba MDC at high radon stations, but increase them at low radon stations
Note sample storage complexity
Increased detector efficiency, background reduction, & increased flow can compensate for the loss in MDC
A Gamma-Gamma Coincidence Detection System for Radionuclide Aerosol Monitoring
A. Couture, A. Carmen, H. Miley, H. Garland, L. Lidey, M. Larche, M. Moore, N. Elmore, R. Guerrero, S. Cheng, and T. Butcher
O3.2-251

## Slide 4
RASA
Radionuclide Aerosol Sampler/Analyzer was developed at PNNL in the 90s
Automated particulate collection, encapsulation, and measurement
A Gamma-Gamma Coincidence Detection System for Radionuclide Aerosol Monitoring
A. Couture, A. Carmen, H. Miley, H. Garland, L. Lidey, M. Larche, M. Moore, N. Elmore, R. Guerrero, S. Cheng, and T. Butcher
O3.2-251
H Miley, et al., (1998) J. Radioanal. Nucl. Chem., 235 (1–2) (1998), pp. 83-87, 10.1007/BF02385942
JL Burnett (2019) Nucl. Inst. Meth. A, Vol. 944 162473 https://doi.org/10.1016/j.nima.2019.162473

## Slide 5
RASA 2.0
A Gamma-Gamma Coincidence Detection System for Radionuclide Aerosol Monitoring
A. Couture, A. Carmen, H. Miley, H. Garland, L. Lidey, M. Larche, M. Moore, N. Elmore, R. Guerrero, S. Cheng, and T. Butcher
O3.2-251
ESP Collector
Sample Packaging
Dual Gamma Detectors
Creare
See Swanwick O3.2-175
PNNL
PNNL/Creare

## Slide 6
Modeling for Detector Selection
Mylar Source: 76 x 76 x 0.63 mm (3” x 3” x ¼”)
Detectors – (single detectors & pairs considered)
  Stock and custom BEGe’s
  XTRa’s – 4 large stock models
  Original RASA (with filter) for baseline comparison
Energy Range – 60-1600 keV
BKGD estimated based solely on 212Pb – efficiency to downscattering into ROIs (focus on140Ba)
Calculation of detector efficiency & relative MDCs
Cost/Sens. FOM = Rel. 140Ba MDC × Price (Small = 👍)
MCNP1 used – no consideration of coincidence
Mirion XTRa GX10020 selected
Capable of 4 samples/day at ~2.5x air flow
A Gamma-Gamma Coincidence Detection System for Radionuclide Aerosol Monitoring
A. Couture, A. Carmen, H. Miley, H. Garland, L. Lidey, M. Larche, M. Moore, N. Elmore, R. Guerrero, S. Cheng, and T. Butcher
O3.2-251
https://www.mirion.com/products/technologies/spectroscopy-scientific-analysis/gamma-spectroscopy/detectors/hpge-detectors-accessories/xtra-extended-range-coaxial-ge-detectors
1C. J. Werner, et al. (2018)., Los Alamos National Laboratory, report LA-UR-18-20808. https://mcnp.lanl.gov/pdf_files/TechReport_2018_LANL_LA-UR-18-20808_WernerBullEtAl.pdf

## Slide 7
Dual Gamma Development Testbed – ARGO2
Advanced Radionuclide Gamma-spectrometer 2 (ARGO2)
Pair of Mirion BE6530 in Pb cave with LYNX2 MCAs
Control software & data processing development
A Gamma-Gamma Coincidence Detection System for Radionuclide Aerosol Monitoring
A. Couture, A. Carmen, H. Miley, H. Garland, L. Lidey, M. Larche, M. Moore, N. Elmore, R. Guerrero, S. Cheng, and T. Butcher
O3.2-251
M. Sharma, et al., (2024) Trans. of the Am. Nucl. Soc., Vol 128, 852-853, DOI:10.13182/T128-41654

## Slide 8
Software Development
C++/Linux control software/WebGUI framework used in PNNL-designed beta/gamma coincidence systems (Xenon International, RL16)
Control server developed for pairs Mirion LYNX MCAs in TLIST mode
Live 8192 channel singles/coincidence histogram generation (20 kB/10s-100s MB)
Updates to deal with high(er) count rate scenarios & detector resolution
  Accidental coincidences
  Coincidence dead times
  Updated coincidence logic
Binary listmode data files available, but not needed for routine operation
~1.5 MB/min or 2GB/day at ~1000 cps
SoH generation – count rates, HV, cooler temperatures/power consumption, leakage currents
“First Look” detector – H3D M400 data acquisition & spectrum viewing
A Gamma-Gamma Coincidence Detection System for Radionuclide Aerosol Monitoring
A. Couture, A. Carmen, H. Miley, H. Garland, L. Lidey, M. Larche, M. Moore, N. Elmore, R. Guerrero, S. Cheng, and T. Butcher
O3.2-251

## Slide 9
Software Development
A Gamma-Gamma Coincidence Detection System for Radionuclide Aerosol Monitoring
A. Couture, A. Carmen, H. Miley, H. Garland, L. Lidey, M. Larche, M. Moore, N. Elmore, R. Guerrero, S. Cheng, and T. Butcher
O3.2-251
Singles Histograms
Coincidence Histograms
Acquisition Control
SoH

## Slide 10
Geant4 Simulation – ARGO2
Simulation of ARGO2 using the Geant4 Cascade Summing Correction application
Model validated with comparison to mixed gamma standard
Near term plan
  Model RASA2.0 geometry
  Efficiency to 140Ba, 131I, 99Mo/99mTc
  Estimate ambient & sample-induced BKGD for all IMS network stations with RASA2.0 → MDCs
  Consider coincidence & Compton-summed detector pairs
  Use ATM models to estimate network sensitivity for varying collection/count times using RASA2.0
A Gamma-Gamma Coincidence Detection System for Radionuclide Aerosol Monitoring
A. Couture, A. Carmen, H. Miley, H. Garland, L. Lidey, M. Larche, M. Moore, N. Elmore, R. Guerrero, S. Cheng, and T. Butcher
O3.2-251
1Pierson, B., Hagen, A., & Archambault, B. (2022) G4CSC, 0.1a. Retrieved from G4CSC - Geant4 Cascade Summing Corrections: https://gitlab.pnnl.gov/ares/g4csc

## Slide 11
Next Steps
Air collection campaign to assess background vs. decay time using ARGO2
Evaluate RASA2.0 sensitivity for extended set of IMS-relevant radionuclides
Evaluate the integration of coincidence data into the IDC spectral analysis methodology
Further software development
  Auto gain stabilization/matching
  Historical SoH & spectal data viewing
  QC-source measurements
  External control
Build a prototype for integration with Creare ESP
  Detector acquisition
  Further mechanical design
A Gamma-Gamma Coincidence Detection System for Radionuclide Aerosol Monitoring
A. Couture, A. Carmen, H. Miley, H. Garland, L. Lidey, M. Larche, M. Moore, N. Elmore, R. Guerrero, S. Cheng, and T. Butcher
O3.2-251

## Slide 12
A Gamma-Gamma Coincidence Detection System for Radionuclide Aerosol Monitoring
A. Couture, A. Carmen, H. Miley, H. Garland, L. Lidey, M. Larche, M. Moore, N. Elmore, R. Guerrero, S. Cheng, and T. Butcher
O3.2-251
Thank you. Questions?