## Slide 1
Authors : H.-D. Lenouvel1, H. Paradis1
Co-authors : A. de Vismes Ott1, P. Gross1, S. Topin1,2
1 CEA, DAM, DIF, F-91297 Arpajon, France
2 CNAM, Analyse Chimique et Bioanalyse, EPN-7, 292 rue Saint Martin, 75003 Paris, France
09 september 2025
Enhancements of FRL08 : metrological connection between simulation/experiment and utilization of AI tools for coincidence matrix analysis
O3.2-267

## Slide 2: From context to modeling : systems and methods

## Slide 3
Resources deployed by the CEA as part of the CTBT framework:
  Stations: SPALAX and SPALAX-NG [1]
  Certified Laboratory: FRL08 (Gamma3) [2]
  National Data Center (NDC)
Research for radionuclide measurements:
  Coincidence measurements (with the SPALAX-NG and Gamma3)
  Data analysis : AI and spectral unmixing methods (P3.2-263)
Radionuclide technology : context and stakes
Fig 1. SPALAX-NG
Fig 2. Gamma3 (left) & PIPSBox (right)
[1] S. Topin et al., J. Environ. Radioact. 225 (2020) 106442
[2] A. Cagniant et al., Applied Radiation and Isotopes 98 (2015)

## Slide 4
Model creation with MCNP
 Use of sources of different sizes and compositions to provide a general calibration
Standard sources:
  Energy range covered: [22 - 2505 keV]
  Geometries:
  compressed or uncompressed particle filters (diameter of 110 mm)
  volumetric geometries (from 20 mL to 500 mL)
Modeling from:
  Technical drawings provided by the manufacturer
  Radiographies of detectors
Fig 3. Radiography of a HPGe detector (BEGe5030P)
Fig 4. Display of a cross-sectional view of a detector modeled under MCNP.
Fig 5. Common sources geometry used by the laboratory
BP20
3M
RASA
PIPSBox geometry
Double Face

## Slide 5
Chart: Energy (keV)
| 5.41 | 5.95 | 11.87 | 13.93 | 14.41 | 17.55 | 22.1 | 24.1 | 25.07 | 26.34 | 27.5 | 32.06 | 33.299999999999997 | 36.450000000000003 | 37.33 | 38.06 | 59.54 | 60 | 60 | 88.03 | 122.06 | 136.47 | 165.86 | 255.13 | 320.08 | 391.7 | 514 | 661.66 | 834.85 | 898.04 | 1115.55 | 1173.23 | 1332.49 | 1836.07 | 2734.09
Mean | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065
Standard deviation (k=2) | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445
+2 sigma | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586
BP20 | 0 | 0 | 0 | 0 | 1.0152 | 0 | 0 | 0 | 0 | 0.96382 | 0 | 0 | 0 | 1.0307 | 1.029 | 0.99988 | 1.0044 | 1.0406 | 0.99959 | 0.9688 | 1.0045 | 1.0088 | 0.98948 | 0.97571 | 0.97131 | 0.9763 | 0.96841 | 0
3M | 0 | 0 | 0 | 0 | 0.88469 | 0 | 0 | 0 | 0 | 1.2171 | 0 | 0 | 0 | 0.97303 | 0.99843 | 1.0145 | 0.96188 | 1.0525 | 1.0134 | 1.0273 | 0 | 0.99888 | 1.0257 | 1.0211 | 0.99725 | 0.98106 | 0.99135 | 0.98879 | 0
Double_Face | 0 | 0 | 0 | 0 | 1.0405 | 0 | 0 | 0 | 0 | 1.0125 | 0 | 0 | 0 | 0 | 1.0199 | 1.0308 | 1.0266 | 1.0044 | 0 | 0.9998 | 1.0247 | 0 | 0.99948 | 0.98876 | 0 | 0 | 0 | 0 | 0 | 0
PIPSBox Geometry | 0 | 0 | 0 | 0 | 0.98805 | 0 | 0 | 0 | 0 | 1.0174 | 0 | 0 | 0 | 1.0335 | 1.0461 | 1.0244 | 0.97873 | 0.98177 | 1.0492 | 0.96621 | 0.96224 | 0 | 0
rasa | 0 | 0 | 0 | 0 | 1.2013 | 0 | 0 | 0 | 0 | 0.96434 | 0 | 0 | 0 | 0.96774 | 0.98676 | 0.98365 | 1.0071 | 1.0003 | 1.0889 | 1.0163 | 0 | 0.99852 | 1.0225 | 0.97426 | 0.99773 | 0.98954 | 0.98195 | 0.97477 | 0
1° Modeling of detectors
2° Optimization:
Mean ratio =
3° Validation of models
60 keV
Parameters to optimize:
  Thickness of materials (dead layers, carbon window…)
  Position of the source relative to the crystal detector
  Crystal size
 Model validation enables simulation for calculating new detection efficiencies
Detection efficiency:
Validation of the digital model:
Fig 7. Detection efficiency ratio (EXP/MCNP) versus energy
Fig 6. Modeling of the configuration (ATHOS + ARAMIS) & PIPSBox
2 conditions must be met for the model:
0.99<Rmean<1.01
<20% below 60 keV
<10% above 60 keV
from A. Berlizov, MCNP-CP Upgrade Patch Version 3.2, Institute for Nuclear Research National Academy of Sciences of Ukraine, 2012
Validation of models

## Slide 6: Quantification through coincidence measurements

## Slide 7
Validation for -ray spectrometry and for coincidence spectrometry. Very good agreement with reference values.
H.-D. Lenouvel, et al., Measurement of radioxenon isotopes for nuclear explosion detection using coincident detector calibrated by simulation, ARI. DOI: ARI 111886
RXe | (Bq) | k=2 (Bq) | (Bq) | k=2 (Bq) | Deviation (%) | score
131mXe | 94 | 10 | 84 | 13 | -10.6 | -0.60
133Xe | 204 | 23 | 205 | 19 | 0.8 | 0.06
133mXe | 10.2 | 1.1 | 9.1 | 1.4 | -11.3 | -0.63
135Xe | 217 | 22 | 226 | 21 | 4.0 | 0.29
RXe |  |  |  |  | Deviation (%) | 
131mXe | 94 | 10 | 84 | 13 | -10.6 | -0.60
133Xe | 204 | 23 | 205 | 19 | 0.8 | 0.06
133mXe | 10.2 | 1.1 | 9.1 | 1.4 | -11.3 | -0.63
135Xe | 217 | 22 | 226 | 21 | 4.0 | 0.29
Measurement:
  (ATHOS+ARAMIS)&PIPSBox
  Duration: 57600 s
Analysis:
Efficiency calculation:
  MCNP-CP
RXe | β/γ coincidence spectrometry
131mXe | 0.16 ± 0.02
133Xe | 0.20 ± 0.02
133mXe | 0.147 ± 0.023
135Xe | 0.176 ± 0.016
Fig 8. matrix obtained from the measurement
Tab 1. detection efficiencies for the 4 radioxenon isotopes
Tab 2. Comparison between reference activities and measured activities
Validation of the method on a gaseous sample

## Slide 8
Identification of 134Cs using coincidences:
Measurement of a fresh fission products sample from CTBTO:
  Classic -ray spectrometry: complex spectra (~300 peaks)
  Coincidence spectrometry: characteristic signatures
High SNR values make it easier to qualify and quantify certain radionuclides
Fig 9. Decay scheme of 134Cs
Fig 10. Particulate filter placed between two HPGe detectors with expected coincidence matrix signatures
Fig 11. Classic 𝛾-ray with zoom around 605 keV
Fig 12. Coincidence matrix with a zoom on the 134Cs
: Signal-to-Noise Ratio
Benefits of using coincidence detection (example with 134Cs)

## Slide 9
Validation of the method on a IAEA sample
Radionuclide | (Bq) | k=2 (Bq) | (Bq) | k=2 (Bq) | Deviation (%) | score
110mAg | 7.51 | 0.09 | 7.36 | 0.70 | -2.0 | -0.21
139Ce | 7.51 | 0.09 | 6.7 | 1.1 | -11.0 | -0.78
134Cs | 7.56 | 0.07 | 7.75 | 0.79 | 2.5 | 0.24
152Eu | 7.55 | 0.07 | 8.0 | 0.8 | 5.7 | 0.55
235U | 0.0747 | 0.0009 | 0.062 | 0.027 | -16.2 | -0.46
Radionuclide |  |  |  |  | Deviation (%) | 
110mAg | 7.51 | 0.09 | 7.36 | 0.70 | -2.0 | -0.21
139Ce | 7.51 | 0.09 | 6.7 | 1.1 | -11.0 | -0.78
134Cs | 7.56 | 0.07 | 7.75 | 0.79 | 2.5 | 0.24
152Eu | 7.55 | 0.07 | 8.0 | 0.8 | 5.7 | 0.55
235U | 0.0747 | 0.0009 | 0.062 | 0.027 | -16.2 | -0.46
Analysis:
Detecting radionuclides (134Cs) with poor visibility in -ray spectrometry
Validation for -ray spectrometry and for coincidence spectrometry. Very good agreement with reference values.
Tab 3. Comparison between reference activities and measured activities
Fig 13. matrix obtained from the measurement with a zoom on the 134Cs
Standard IAEA swipe sample
Measurement:
  ATHOS&ARAMIS
  Duration: 316800 s
ROI
Bkgd
Efficiency calculation:
  MCNP-CP

## Slide 10: Qualitative AI model

## Slide 11
Memory-intensive 2D spectrum  spectra size reduced by ~300x
Composition of simulated samples : ~130 000 spectra
  ~3% : spectral signature of each radionuclide of the relevant list
  ~18% : fission products from 235U (High Enriched Uranium) with fast neutrons (14 MeV)
  ~79% : random distributions
Convolutional Neural Network (CNN):
Database & CNN architecture
Ensuring both physical accuracy and sufficient representation of each radionuclide
Setting thresholds for each radionuclide to determine their presence
Fig 14. Representation of the CNN architecture employed

## Slide 12
Evaluation of the spectral CNN
Determine the model’s ability to discriminate between positive and negative classes :
Precision-Recall (PR) curve:
High detection accuracy for specific classes
Exclusion of certain classes treated as noise
Determine the best threshold
F1-score (harmonic mean of the precision and recall):
For this model:
Precision: 0.78
Recall: 0.84
F1-score: 0.79
Fig 16. PR curve for 136Cs
Fig 15. Distribution of positives/negatives as a function of the raw output
Threshold

## Slide 13
Radionuclides in the source:
  Detectable by the model : 57Co, 60Co, 88Y, and 139Ce
  Undetectable by the model : 54Mn, 65Zn, 109Cd, 113Sn, 137Cs, and 241Am
CNN results:
  57Co : only visible at low energies
Results from a standard source
Radionuclide | A (Bq) | Threshold (%) | Predicted probability of presence (%)
60Co | 185.5 | 69 | 85
139Ce | 919.5 | 73 | 86
88Y | 2794 | 70 | 81
Radionuclide | A (Bq) | Threshold (%) | Predicted probability of presence (%)
57Co | 105.2 | 71 | 24
Detected
Not detected
Most of the elements detectable by the model are correctly classified
No false positives were observed
Tab 4. Model results on a 10-element standard source

## Slide 14
CNN results:
Results from the PTE 2023 (3M geometry)
Radionuclide | Importance | A (Bq) | Threshold (%) | Predicted probability of presence (%)
125Sb | major | 2.8 | 74 | 78
136Cs | major | 21.2 | 66 | 82
95Nb | major | 26.2 | 62 | 85
147Nd | major | 35.1 | 70 | 74
95Zr | major | 63.7 | 69 | 96
103Ru | major | 81.0 | 70 | 98
140Ba | major | 93.8 | 70 | 93
140La | major | 108.0 | 71 | 93
106Ru | / | / | 68 | 82
124Sb | major | 2.1 | 68 | 39
156Eu | minor | 3.6 | 71 | 43
132Te | minor | 7.0 | 71 | 16
127Sb | minor | 10.2 | 72 | 41
126Sb | major | 17.5 | 65 | 35
131I | major | 27.6 | 69 | 63
134Cs | / | / | 66 | 19
Detected
Not detected
Most of the elements detectable by the model are correctly classified
Most undetected elements are due to low activity/emission levels or insignificant coincidence signatures
Tab 5. Model results on a 25-element sample
106Ru
134Cs
Fig 17. Coincidence matrix with a zoom on the 106Ru and 134Cs

## Slide 15: Conclusion & Perspectives

## Slide 16
Conclusion
Metrological connection with Monte Carlo simulation :
    Model optimization and validation with standard sources
    Determination of detection efficiencies :
      Validation with a radioxenon gaseous sample ( coincidence)
      Validation with IAEA swipe intercomparison exercise ( coincidence)
Implementation of a first AI model for matrix analysis :
    Feature detection with a CNN
    Qualitative multi-label classification
Thank you.
Optimization of the AI tool:
    Refine the dataset, improve the network structure and the loss function, for radionuclides whose detection is challenging
    Include uncertainties to the predicted probabilities of presence
Explore the use of spectral unmixing method for coincidence measurements
    Adapt the tools developed by C.-P. Mano (P3.2-263) to the context of HPGe applications
Perspectives
RXe | (Bq) | k=2 (Bq) | (Bq) | k=2 (Bq) | score
131mXe | 94 | 10 | 84 | 13 | -0.60
133Xe | 204 | 23 | 205 | 19 | 0.06
133mXe | 10.2 | 1.1 | 9.1 | 1.4 | -0.63
135Xe | 217 | 22 | 226 | 21 | 0.29
RXe |  |  |  |  | 
131mXe | 94 | 10 | 84 | 13 | -0.60
133Xe | 204 | 23 | 205 | 19 | 0.06
133mXe | 10.2 | 1.1 | 9.1 | 1.4 | -0.63
135Xe | 217 | 22 | 226 | 21 | 0.29

## Slide 17: Annexes

## Slide 18
Gamma3 : configuration for the measurement of gaseous samples
Configuration:
  2 HPGe: detection of photons ()
  2 PIPS: detection of electrons ()
  The PIPSBoxTM is placed between ATHOS and ARAMIS detectors
coincidences:
Fig 10. (ATHOS+ARAMIS)&PIPSBox configuration
Fig 11. Depiction of the shape of the fingerprints
Fig 13. particle emission energy spectrum
Fig 12. Energy spectrum of the electron emitted by internal conversion

## Slide 19
Benefit of using coincidence detection
The benefit of using coincidences in the X-ray region:
Characteristic signatures of radionuclide (very low background)
Fig 15. -ray spectrum with a zoom on the X-ray region
Fig 14. Separation of the signatures of metastable radioxenon isotopes

## Slide 20
Radionuclide | spectrometry : DL (mBq) | spectrometry : DL (mBq)
131mXe | 73 | 0.3
133Xe | 3.3 | 1.0
133mXe | 33 | 0.2
135Xe | 18.3 | 4.3
Radionuclide |  | 
131mXe | 73 | 0.3
133Xe | 3.3 | 1.0
133mXe | 33 | 0.2
135Xe | 18.3 | 4.3
Radionuclide | spectrometry : DT / DL (mBq) | spectrometry : DT / DL (mBq)
108mAg | 2.03 / 4.2 | 0.39 / 2.9
110mAg | 2.44 / 5.0 | 3.67 / 26.4
60Co | 1.81 / 3.8 | 0.60 / 4.3
125Sb | 5.31 / 10.9 | 5.39 / 19.0
134Cs | 1.75 / 3.6 | 0.71 / 3.2
Radionuclide |  | 
108mAg | 2.03 / 4.2 | 0.39 / 2.9
110mAg | 2.44 / 5.0 | 3.67 / 26.4
60Co | 1.81 / 3.8 | 0.60 / 4.3
125Sb | 5.31 / 10.9 | 5.39 / 19.0
134Cs | 1.75 / 3.6 | 0.71 / 3.2
Calculation of decision thresholds (DT) and detection limits (DL) on a blank measurement:
  ISO standard 11929
  Comparison between the classical and coincidence spectrometry
Particulate radionuclides:
Noble gas sample:
Regarding certain radionuclides :
very low 
 often highly favorable better sample characterization
spectrometry highly favorable for metastable radioxenon isotopes
Tab 6. Comparison of DT/DL between and spectrometry for particulate measurements
Tab 7. Comparison of DT/DL between and spectrometry for noble gas measurements

## Slide 21
Training of the spectral CNN
Activation function : sigmoid  raw outputs  probabilities of presence
  multi-labels classification
Loss function : Custom FocalLoss [5] with BCEWithLogit

   balances positive vs. negative classes
   focuses on hard vs. easy examples
Optimizer : Adam (step : )
[5] T.-Y. Lin et al., Facebook AI Research (FAIR), 2017
Here :
Epochs : 15
Batch : 128 spectra
/
Strong class imbalance + avoiding underfitting of difficult classes